## Slide 1
Balloon Occlusion
Robbie Lendrum
Deputy Clinical Lead and Training Lead
Jan Jansen
Clinical Chief Investigator

## Slide 2: Open “PACK 3 (OCCLUSION)”

## Slide 3: Identify: ER-REBOA catheter pack 20 ml 0.9% Saline Drawing up Needle 3 way tap 10ml syringe

## Slide 4: Select occlusion zone

## Slide 5: Zone III (aortic bifurcation): Isolated pelvic injuries Junctional injuries

## Slide 6: Zone I (thoracic aorta): Abdominal (+/- pelvic) injuries

## Slide 7: Measure

## Slide 8: Surface markings for zone III - Umbilicus

## Slide 9: Surface markings for zone I - Mid-Sternum

## Slide 11: Zone I: Approximately 46 cm Zone III: Approximately 28 cm

## Slide 12: Attach transducer to “ART” lumen 3-way tap Attach inflation syringe (10 ml saline) to “BAL” using additional 3-way tap

## Slide 13: Wet Balloon

## Slide 14: Slide orange sheath over P-tip to straighten

## Slide 15: Flush catheter

## Slide 16: Insert assembly into arterial sheath 0.5cm and advance catheter only approximately 15cm

## Slide 17: Withdraw and peel away orange sheath

## Slide 18: Insert catheter to desired depth (zone I or zone III)

## Slide 19: Inflate balloon with saline

## Slide 20: Suggested inflation starting volumes: Zone III: 2 ml Zone I: 8 ml

## Slide 21: You may need more – gauge against haemodynamic response

## Slide 22: But beware of overinflation – experience has shown this to be very common

## Slide 23: Turn “BAL” stopcock off to balloon

## Slide 24: Record time of balloon inflation

## Slide 25: Open “PACK 4 (SECURING)”

## Slide 27: Secure catheter to prevent migration

## Slide 28: Quickly progress to definitive haemorrhage control