## Slide 1
Who is www.ZoomAFM.com ?
Song Xu Application scientist of Agilent Technology Inc
Steve Raba Service engineer of Agilent Technology Inc
XQ Tong Technical Manager of Agilent Technology
Gerald Kada Application Scientist Germany
Louis Pacheco Former Application Engineer France, Founder of CSInstrument
… … … (there will be more to join us, from around the world)
Why we are here ?
Keep your instrument running for 10 more years.
Support Molecular Imaging, Agilent and Keysight AFM users.
Provide lectures and training
Repair and maintenance
Our purpose: we want to provide a consistent support with easy to find communication. Welcome to ask questions or just have a chat, and please be prepared to introduce yourself and what you do with your instrument.

## Slide 3
Basic Physics of Atomic Force Microscope III

## Slide 4
Basics of AC / tapping mode
  Basic force interaction involved in tapping mode imaging
  Tappping mode, true non-contact mode, dipping mode
  Understand how to estimate and measure tapping amptlitude
  How to estimate tapping force changing with setpoint and driving power
  Basic steps to achieve high resolution with AC mode.
Content:

## Slide 5
Force Interaction of AC / Tapping Mode

## Slide 6
Tapping / AC / non-contact?
Tapping: higher amplitude to ensure short range repulsive interaction
Non-contact: small amplitude interaction at attractive range distance

## Slide 7
Sliding friction or cycling shear
Adhesive debonding
Cycling compression/tension/debonding
Motion
Force on tip
Analysis of basic tip sample interaction and energy dissipation
Each of these modes can be used to contrast materials and analyze viscoelastic properties.

## Slide 8
AC / Tapping Mode Imaging on Close Range Force Interaction
< 10 nm
<0 nm

## Slide 9
Cantilever tip: a simple harmonic oscillator
    ½,

## Slide 10
When the tip interacts with sample

## Slide 11
Driving
X scan
AFM cantilever tip as driven harmonic oscillator
=
    ½,
K0 b
m

## Slide 12
What happen when tip interacts with sample

## Slide 13
Separation distance
No interaction
Force const ( K = df / ds )
Repulsive region (contact / tapping)
Positive force gradient
Attractive force region (non-contact)
Negative force gradient
Cantilever force constance
Net force constance
Separation distance
Attractive region: resonance f shift lower
Repulsive region: resonance f shift higher
Resonance frequency affected by force gradient: a simple harmonic oscillator
    ½,

## Slide 14
Force (positive = force vector from sample surface toward tip)
Tip
sample
Separation distance
AFM tip experience two different forces:
Force interaction from the sample
Deflection force from cantilever
Separation distance
No interaction
Net force
Tip-sample interaction
Cantilever deflection force
Force const ( K = df / ds )
Repulsive region (intermittened contact / tapping)
Attractive force region (non-contact)
Negative force gradient
Cantilever force constance
Net force constance
Separation distance
Resonance frequency affected by force gradient: a simple harmonic oscillator

## Slide 15
Amplitude (nm)
Phase (degrees)
Frequency (kHz)
Net Repulsive
Driving at resonance
A0 = Afree
A
f
Amplitude (nm)
Phase (degrees)
Frequency (kHz)
Net Attractive
F0
Driving at resonance F0
Qcant = F0/DF ≈ 300
DF
How force gradient changes resonance frequency and phase

## Slide 16
z
d
approach
widthdraw
approach
How AFM cantilever tip detects sample—AC mode

## Slide 17
Basic AFM imaging mode—contact mode & AC mode
Repulsive region (intermittened contact / tapping)
Repulsive region (intermittened contact / tapping)
Attractive force region (non-contact)
Actual topo
Amplitude
error
Feedback output

## Slide 18
Phase

## Slide 19
Thank you for your attention
To be continued
Next week:
Basic to the basic, look at a life instrument
Life Demo!
Every Tuesday morning
support@ZoomAFM.com
Every Tuesday morning 11:00 am EST
Consider add: Monday afternoon 3:00 pm EST ??