Sterile Processing Department: Log tracking guide v 8/22/2025

| Log | Question | Answers | Steps |
| --- | --- | --- | --- |
| Autoclave Steam Sterilizers | Total loads history | All instances of uses | Management; reporting; run Reports; Monthly Load Summary, choose dates and specific sterilizer, “Loads” is the total number of uses; also gives Bowie dick info |
|  | Daily Biological Indicator test | Bowie Dick test done before the sterilizer is used for the day | Management; reprocessing; Load History; select sterilizer/washer to be viewed; select date range; in the Cycle column, click on the Pinpoint and select Bowie Dick or IUSS Spore Test |
|  | Individual load testing | Biological indicator test done on each load sterilized | Management; reprocessing; Load History; select sterilizer/washer to be viewed; select date range; in the Cycle column, ensure all options are chosen; review the “BIs” column, Green checkmarks indicate successful BI load status |
|  | Biological Indicator validation test | Test done to ensure the BI used in individual loads are valid | BI can be found in “test Results” report. Management>reports>run reports> test results (select date range, recommend this done daily or weekly) and will show all tests in on section. This will have bios and leaks! |
|  | Leak test | Test done weekly | Management; reprocessing; Load History; select sterilizer/washer to be viewed; select date range; in the Cycle column, click on the Pinpoint and select Leak |
| Low temperature sterilizers (STERRAD, VPro) | Individual load testing | Biological indicator test | Management; reprocessing; Load History; select sterilizer/washer to be viewed; select date range; in the Cycle column, ensure all options are chosen; review the “BIs” column, Green checkmarks indicate successful BI load status |
|  | Biological Indicator validation test | Test done to ensure the BI used in individual loads are valid | BI can be found in “test Results” report. Management>reports>run reports> test results (select date range, recommend this done daily or weekly) and will show all tests in on section. This will have bios and leaks! |
|  | Leak test | Test done weekly | Management; reprocessing; Load History; select sterilizer/washer to be viewed; select date range; in the Cycle column, click on the Pinpoint and select Leak *STERRAD Leak test is incorporated into the regular cycle, so no Leak test needs to be run |
| High level disinfectors | Individual load testing | Test strip indicators done on each load sterilized | Download data from Medivator until data is available in SPM |
|  | Test strip validation | Test strips are checked for validation for use | Review strip bottles against file of certifications; for MEdivators: Certificates of Analysis (COA) ¦ STERIS |
| Instrument / Cart washers | Washing process indicator test | Done on each washer daily | Management Reporting; run Reports; Test Results option from “All Reports”, Choose date ranges, Choose facility, choose process points (washers), filter by process points (specific washer). |
| Ultrasonic washers | Cleaning indicator daily test | Test done daily to ensure cleaning efficacy | Management; reprocessing; Load History; select sterilizer/washer to be viewed; select date range; review Status for “built” (incomplete) loads and correct the record |
| Decontamination sink; Enzymatic Dosing system (AccusInQ) | Daily testing | To include detergent dose and water temperature | Management; reprocessing; Load History; select sterilizer/washer to be viewed; select date range; a daily test is done each morning, clicking on the individual date will identify the Test Results screen identifying “Pass”. |
| Endoscope flushing system (Scope Buddy) | Daily testing performed? | Flow validation test is to be done each day before use | Management; reprocessing; Load History; select Scope Buddy to be viewed; select date range; in the Cycle column, select “Scope Flow Validation” a completed test should be done each day the Scope Buddy is used. |
| Eyewash station log | Weekly testing performed | Testing to be charted weekly | Current tracked via paper on specific eyewash station, by Facilities |
| Drying cabinet | No daily testing expectation; cleaning schedule should be in place | MIFUs identify weekly cleaning expectation on most drying cabinets | Currently tracked via paper in dept |
| Room temperature & humidity logs | Daily testing | Charted daily | Current tracked electronically by Facilities |
| Positive / Negative air pressure logs | Daily testing | Charted daily | Current tracked electronically by Facilities |