EFFECT OF HME ON AEROSOL DRUG DELIVERY AND AIRWAY RESISTANCE

IN SIMULATED VENTILATOR DEPENDENT ADULTS USING JET AND MESH NEBULIZERS

Arzu Ari, Abdulrahman Alkhathami, Rowaida Qoutah, Ahmad Almmary, James B. Fink

Georgia State University, Department of Respiratory Therapy, Atlanta, GA, USA

BACKGROUND: Placement of a heat moisture exchanger (HME) between aerosol generator and patient has been associated with greatly reduced drug delivery and increased resistance of gas passing through the HME increasing  work of breathing. The purpose of this study was to evaluate the effect of  a specific HME placed between nebulizer and patient on aerosol deposition and airway resistance (Raw) in simulated ventilator dependent adults.

METHODS: An in vitro lung model was developed to simulate a mechanically ventilated adult (Vt 500 ml, RR 15/min, PEEP 5 cmH2O and I:E ratio: 1:2) to an intubated adullt manikin with an endotracheal tube (8 mmID). The  mainstem bronchi of the manikin was connected to a Y adapter through a collecting filter (Respirgard II) attached to a test lung through a heated cascade humidifier (37⁰C producing 100% relative humidity) to simulate exhaled humidity. For treatment conditions, an HME (ThermoFloTM 6070, ARC Medical) was placed between the ventilator circuit at the ETT and allowed to acclimate to the exhaled heat and humidity for 30 min prior to aerosol administration. The values on airway resistance (cmH20/l/s) was taken from the display on the ventilator monitor (Hamilton Galileo) at 0, 10, 20 and 30 min after HME placement and after each  of 4 aerosol treatments. Albuterol sulfate (2.5 mg/3mL) was administered with jet (MistyMax10, Airlife) and mesh (Aerogen Solo, Aerogen) nebulizers positioned in the inspiratory limb 6 in from the Y adaptor and at the Y, respectively. Control consisted of nebulization with no HME. Drug was eluted from filter at the end of the trachea and measured using spectrophotometry (276 nm).

RESULTS: The table shows mean±SD percent dose delivered and Raw.  Greater than 60% of the control dose was delivered through the ThermoFloTM. No significant difference was found between the first four treatments given by the jet (p=0.825) and the mesh (p=0.753) nebulizer. There was a small but significant increase in Raw between pre- and post-4 treatments with the jet (p=0.001) and mesh (p=0.015) nebulizers.

CONCLUSION: The ThermoFloTM HME effectively passed the majority of aerosol on to the airway. Increases in Raw would likely not be outside of a tolerable range in ventilated patients.  Further research with other HMEs and materials is warranted.

|  | TREATMENT GROUPS (WITH HME) |  | CONTROL GROUPS (WITHOUT HME) |  |
| --- | --- | --- | --- | --- |
| Measuring Aerosol Deposition | Jet Nebulizer | Mesh Nebulizer | Jet Nebulizer | Mesh Nebulizer |
| 1st treatment | 3.47 ± 0.40 | 6.61 ± 0.34 | 5.44 ± 0.17* | 10.64 ± 0.53* |
| 2nd treatment | 3.56 ± 0.52 | 6.54 ± 0.71 |  |  |
| 3rd treatment | 3.38 ± 0.41 | 6.57 ± 0.84 |  |  |
| 4th treatment | 3.64 ± 0.44 | 6.40 ± 0.66 |  |  |
| Measuring Raw (cmH20/l/s) |  |  |  |  |
| Before treatment | 15.50 ± 0.57 | 13.75 ± 0.50 |  |  |
| After treatment | 17.50 ± 0.57 | 16.75 ± 1.70 |  |  |

* Delivered dose greater than treatment with same nebulizer (p=0.001)