## 20 cP 200 mmHg tests
Bronchoscope Model | Insertion Cord Diameter (mm) | Min Working Channel Diameter (mm) | Trial 1 (g) | Trial 2 (g) | Trial 3 (g) | Trial 4 (g) | Trial 5 (g) | Mean (g) | STDEV | Viscosity (cP) | Working Channel Size (mm)
Brand A Slim (3.8/1.2) | 3.8 | 1.2 | 4.6 | 4.6 | 4.5 | 4.4 | 4.3 | 4.48 | 0.13038 | 20.03 | 1.2
Brand B Slim (3.8/1.2) | 3.8 | 1.2 | 5.9 | 6 | 5.9 | 6 | 5.9 | 5.94 | 0.054772 | 20.03 | 1.2
Brand C Slim (3.2/1.2) | 3.2 | 1.2 | 3.8 | 3.5 | 3.7 | 3.6 | 3.5 | 3.62 | 0.13038 | 20.03 | 1.2
Bronchoscopus Rex™ Built-in Slim (3.8/1.7) | 3.8 | 1.4 | 28.3 | 28.5 | 28.3 | 28.4 | 28.4 | 28.38 | 0.083666 | 20.03 | 1.7
Bronchoscopus Rex™ Std Slim (3.8/1.7) | 3.8 | 1.4 | 32.7 | 26.9 | 27.2 | 26.8 | 26.4 | 28 | 2.6429 | 20.03 | 1.7
Incomplete Bronchoscopus Rex™ Std. Slim (Additional Tech) (3.9/NA) | 3.9 | 2.0 | 71.3 | 79.9 | 67.1 | 63.7 | 59.5 | 68.3 | 7.8038 | 20.03 | NA
Brand A Slim
Brand B Slim
Brand C Slim
Notes
suction tested between each bronchoscope and each suction container emptying
1000 g Water 2.2 gm Guar Gum
Actual Viscosity: 20.03 cP
Working channel and suction tube cleared with suction of air between each trial
Rotor 0
RPM 12

## 531 cP 200 mmHg tests
Bronchoscope Model | Insertion Cord Diameter (mm) | Min Working Channel Diameter (mm) | Trial 1 (g) | Trial 2 (g) | Trial 3 (g) | Trial 4 (g) | Trial 5 (g) | Mean (g) | STDEV | Viscosity (cP) | Working Channel ID (mm)
Brand A Regular (5.0/2.2) | 5.0 | 2.0 | 4.6 | 4.2 | 4.1 | 4.1 | 4.1 | 4.22 | 0.21679 | 531.1 | 2.2
Brand B Regular (5.0/2.2) | 5.0 | 2.1 | 3.1 | 2.2 | 2.9 | 3.2 | 2.8 | 2.84 | 0.39115 | 531.1 | 2.2
Brand C Regular (4.9/2.2) | 4.9 | 2.2 | 3.5 | 3.5 | 3.5 | 3.5 | 3.4 | 3.48 | 0.044721 | 531.1 | 2.2 | Rotor 3 RPM 60
Rotor 3 RPM 60
Bronchoscopus Rex™ Built-in Regular (5.0/2.8) | 5.0 | 2.5 | 25.5 | 22.5 | 23.3 | 21.8 | 21.9 | 23 | 1.5199 | 531.1 | 2.8
Bronchoscopus Rex™ Std Regular (5.0/2.8) | 5.0 | 2.5 | 17.1 | 16.1 | 14.5 | 15.6 | 15.8 | 15.82 | 0.93648 | 531.1 | 2.8
531.1
Brand A Large (5.8/2.8) | 5.8 | 2.6 | 15.6 | 14.8 | 14.1 | 14 | 13.9 | 14.48 | 0.71903 | 531.1 | 2.8
Brand B Large (5.8/3.0) | 5.8 | 3.0 | 19.8 | 17.5 | 18.1 | 18 | 18 | 18.28 | 0.88148 | 531.1 | 3
Brand C Large (5.8/2.8) | 5.8 | 2.8 | 9.6 | 9.6 | 7.2 | 9.6 | 10 | 9.2 | 1.1314 | 531.1 | 2.8
Bronchoscopus Rex™ Built-in Large (5.8/3.5) | 5.8 | 3.2 | 78.4 | 72.5 | 81.5 | 73.5 | 74.4 | 76.06 | 3.7766 | 531.1 | 3.5
Bronchoscopus Rex™ Std Large (Additional Tech) (5.8/NA) | 5.8 | 3.9 | 41.7 | 38.5 | 38.1 | 37.5 | 38.5 | 38.86 | 1.6395 | 531.1 | NA
Notes
suction tested between each bronchoscope and each suction container emptying
Scopes working channel cleared with water inbetween scope change
Working channel and suction tube cleared with suction of air between each trial
1000.0 g water
5.2 g guar gum