## Slide 1
APD
APD
APD
APD
SH Sensor
Design of a Steerable, Configurable Detection, Adaptive Optics Scanning Laser Ophthalmoscope with Integrated Fixation and Stimulation
S.A. Burns, A. de Castro, L. Sawides. T. Luo, K. Sapoznik, R. L. Warner, Indiana University School of Optometry, 800 E. Atwater, Bloomington, IN, staburns@indiana.edu
We designed, built and tested a near IR Adaptive Optics Scanning Laser Ophthalmoscope system.
Goals for clinical Imaging:
image multiple retinal regions spanning up to 30 degrees with of minimal distortion or skew (montaging and eye tracking accuracy)
Provide flexibility in imaging using both directly backscattered and multiply scattered light image
Be able to measure blood flow based on motion of individual red blood cells.
Provide a fixation channel that can provide visual stimuli and visual stimulation.
System Implementation
Introduction
Discussion
Imaging performance was superb. Multiply scattered light images and confocal images could all be simultaneously collected, with the second optical channel providing offset imaging on a second channel
Results
The system was designed in ZEMAX, then moved to Solidworks for mechanical modeling and finally implemented
Included a steerable system [1,2] for active retinal tracking of more than small eye movements with a large isoplanatic angle to minimize system aberrations when the eye moves.
Flexible detection modalities, including confocal and a Programmable Aperture system to detect multiply scattered [3] light, enhance the visibility of other retinal features.
Dual Offset Imaging to allow precise measurements of temporally varying events.
We designed, built and tested a near IR Adaptive Optics Scanning Laser Ophthalmoscope system. The system performed well in practice, providing superb, flexible imaging and allowing high quality retinal imaging in normal and patients
References
Stephen A. Burns, R. Tumbar, AE Elsner, D Ferguson, & D.X. Hammer Large-field-of-view, modular, stabilized, adaptive-optics-based scanning laser ophthalmoscope. JOSA A, 24 (5), 1313-1326 (2007).
Weyao Zou, X Qi, S.A. Burns, Woofer-tweeter adaptive optics scanning laser ophthalmoscopic imaging based on Lagrange-multiplier damped least-squares algorithm, Biomedical Optics Express, 1986-2004, (2011)
Methods

## Slide 2
APD
APD
APD
APD
SH Sensor
Optics- illumination and steering
DM 1
DM 2
Fast Scan
Galvo 1 Horizontal Position
Galvo 2
Vertical Position
Slow Scan
Detection System
Light Source Supercontinuum Laser
SH Sensor- 30 Hz
Fast Scan 15.8 KHz
Visual Stimulus- TI 4600 display
The system provides steering over 30 degrees. At the design defocus (-2.5D) the isoplanatic angle was greater than 8 degrees. Angular rotation across the field was approximately 3 degrees (along the axis of maximum distortion). This could be compensated by driving both steering mirrors for the slow scan (not implemented).
Back
+10 Degree
+3 Degrees via Fast Scan
Outer boxes- effect of steering to corners of 10 degree field
Center the variation of point spread function for 3 degree scan, which goes through more optics
Optical Design: performance
System corrected for
2.5 diopter Myope at deformable mirror 2.
Vertical Position
Slow Scan
Horizontal Position
Eye
Optical Design Field Positioning
Back
Design of a Steerable, Configurable Detection, Adaptive Optics Scanning Laser Ophthalmoscope with Integrated Fixation and Stimulation
S.A. Burns, A. de Castro, L. Sawides. T. Luo, K. Sapoznik, R. L. Warner, Indiana University School of Optometry, 800 E. Atwater, Bloomington, IN, staburns@indiana.edu
R.Daniel. Ferguson, Z. Zhong, D.X. Hammer, M. Mujat AH Patel. C. Deng, W. Zou, W, S.A. Burns, "Adaptive optics SLO with integrated wide-field retinal imaging and tracking" JOSA A 27, 265-277 (2010)

## Slide 3
APD
APD
APD
APD
SH Sensor
The system provides steering over 30 degrees. At the design defocus (-2.5D) the isoplanatic angle was greater than 8 degrees. Angular rotation across the field was approximately 3 degrees (along the axis of maximum distortion). This could be compensated by driving both steering mirrors for the slow scan (not implemented).
Back
Display System
Optical Design: Implementation
Optical Implementation
Design of a Steerable, Configurable Detection, Adaptive Optics Scanning Laser Ophthalmoscope with Integrated Fixation and Stimulation
S.A. Burns, A. de Castro, L. Sawides. T. Luo, K. Sapoznik, R. L. Warner, Indiana University School of Optometry, 800 E. Atwater, Bloomington, IN, staburns@indiana.edu
Back

## Slide 4
APD
APD
APD
APD
SH Sensor
Programmable Array Detection
Multiply Scattered Light in the retina returns from secondary scatterers after first being refracted and diffracted from retinal structures. Because the actual retinal structure controls the direction of scattering as well as the angle, it is useful to be able to capture different aspects of that scattered light
Schematic of Channel 1 Detection Scheme
A single beam from the retina is sent to three detectors simultaneously. The psf is separated into directly backscattered light by a pinhole at the conjugate plane. The surrounds of the psf (multiply scattered light) is re-imaged onto a DLP, which diverts it to one of two detectors
Confocal Pinhole
Detector 1
Light From Retina
ON
Detector 2
DMD
OFF
Detector 3
Back
Ratio images of the same vessel with different DMD aperture configurations
100 μm
The computed contrast between the “on” and “off” states for the same retinal blood vessel, using different spatial masks. Note that the edges of the vessels are differentially emphasized with mask size and orientation
Back
Design of a Steerable, Configurable Detection, Adaptive Optics Scanning Laser Ophthalmoscope with Integrated Fixation and Stimulation
S.A. Burns, A. de Castro, L. Sawides. T. Luo, K. Sapoznik, R. L. Warner, Indiana University School of Optometry, 800 E. Atwater, Bloomington, IN, staburns@indiana.edu
Toco Y.P. Chui, D.A. VanNasdale and S.A. Burns The use of forward scatter to improve retinal vascular imaging with an adaptive optics scanning laser ophthalmoscope. Biomed. Opt. Express 3(10): 2537-2549. (2012).

## Slide 5
APD
APD
APD
APD
SH Sensor
1. Introduction
Average Velocity: 2.0043
We designed a near IR Adaptive Optics Scanning Laser Ophthalmoscope system with the goal of meeting several requirement for clinical retinal imaging.
the need to image multiple retinal regions spanning up to 30 degrees with of minimal distortion or skew for montaging or eye tracking
A steerable system [1,2] for active retinal tracking of more than small eye movements with a large isoplanatic angle to minimize system aberrations when the eye moves.
Flexible detection modalities, including confocal and a Programmable Aperture system to detect multiply scattered [3] light, enhance the visibility of other retinal features.
Dual Offset Imaging to allow precise measurements of temporally varying events.
Back
Back
Design of a Steerable, Configurable Detection, Adaptive Optics Scanning Laser Ophthalmoscope with Integrated Fixation and Stimulation
S.A. Burns, A. de Castro, L. Sawides. T. Luo, K. Sapoznik, R. L. Warner, Indiana University School of Optometry, 800 E. Atwater, Bloomington, IN, staburns@indiana.edu
Channel 1
Channel 2
Channel 1 is set to use as large displaced aperture identically to channel 2. The two channels enter the pupil at slightly different angles, which produces a displacement. Thus a feature scanned in A at time t, is scanned in channel 2 later (time t+d)
d
Channels are then processed to show only the change over time for each pixel
And pixels exceeding a given z-score in a given capillary are used to detect erythrocytes.
The regions around detected cells are then average in each channel. For the detection channel, this means there is always an RBC in the center
The regions around detected cells are then average in each channel. For the other channel the RBC is not in the center due to blood flow during the time interval. This distance changes with the heartbeat.
de Castro A, Huang G, Sawides L, Luo T, Burns SA. Rapid high resolution imaging with a dual-channel scanning technique. Optics Letters 2016;41:1881-1884.

## Slide 6
APD
APD
APD
APD
SH Sensor
2. System
1. Introduction
Confocal image of cone mosaic. Not cones down to the central fovea are visible.
Scale bar 50 mm.
Gathered in 10 minutes subject time.
X
Back
Design of a Steerable, Configurable Detection, Adaptive Optics Scanning Laser Ophthalmoscope with Integrated Fixation and Stimulation
S.A. Burns, A. de Castro, L. Sawides. T. Luo, K. Sapoznik, R. L. Warner, Indiana University School of Optometry, 800 E. Atwater, Bloomington, IN, staburns@indiana.edu
Montages could be made without the need for adding a skew to the positioning system (frame edges are parallel)
.

## Slide 7
X
Simultaneously collected Image types.
A: Confocal image.
B: Offset Aperture Image from Channel 2
C: Flow map from variance of offset image
D: Image obtained with a Half annular aperture.
E: complement (rest of light) that is not directed to
F: Contrast between D and E (D-E)/(D+E)
Scale bar 100 mm.
A
B
C
D
E
F
Back
Design of a Steerable, Configurable Detection, Adaptive Optics Scanning Laser Ophthalmoscope with Integrated Fixation and Stimulation
S.A. Burns, A. de Castro, L. Sawides. T. Luo, K. Sapoznik, R. L. Warner, Indiana University School of Optometry, 800 E. Atwater, Bloomington, IN, staburns@indiana.edu
Example of a blood vessel image- showing all images extracted from a single 100 frame acquisition. (videos not shown)
.

## Slide 8
X
Simultaneously collected Image of Blood vessel showing the ability to enhance visibility of the vascular walls with a half annular aperture
.
A
B
C
D
E
F
A
B
D
E
F
Simultaneously collected Image types.
A: Confocal image.
B: Offset Aperture Image from Channel 2
D: Image obtained with a Half annular aperture.
E: complement (rest of light) that is not directed to
F: Contrast between D and E (D-E)/(D+E)
Scale bar 100 mm.
Back
Design of a Steerable, Configurable Detection, Adaptive Optics Scanning Laser Ophthalmoscope with Integrated Fixation and Stimulation
S.A. Burns, A. de Castro, L. Sawides. T. Luo, K. Sapoznik, R. L. Warner, Indiana University School of Optometry, 800 E. Atwater, Bloomington, IN, staburns@indiana.edu
D
E

## Slide 9
X
Small Microaneurysm in diabetic, showing confocal (A), offset aperture multiply scattered light(B) and programmable aperture contrast images where presumed endothelials cells in the superficial wall are readily seen ( C)
A
B
C
D
E
F
Confocal foveal cone images (D), simultaneously collected multiply scattered light images showing RPE cells (C ), and shaped aperture images showing that there is a mix of cone and RPE cell contributions at short scattering distances.
Scale bar 100 mm.
Back
Design of a Steerable, Configurable Detection, Adaptive Optics Scanning Laser Ophthalmoscope with Integrated Fixation and Stimulation
S.A. Burns, A. de Castro, L. Sawides. T. Luo, K. Sapoznik, R. L. Warner, Indiana University School of Optometry, 800 E. Atwater, Bloomington, IN, staburns@indiana.edu